## General Contributors-Chemistry
Common Contributors to Measurement Uncertainty – Chemical Analyses (representative list - may not be all inclusive for all types of analyses) (QC sample types in this list are typical of those utilized in AIHA-LAP, LLC laboratories) See additional tabbed sheets for examples
Contributors to Uncertainty | Representative and Applicable QC Data | Comments to Clarify Contributor Effects
Transportation/Storage/Handling:
shipping time, container and temperature | NA / (FS) | NOTE: This is not part of analytical uncertainty, but must be considered by labs providing sampling and when providing guidance regarding sample packaging and shipping. Usually no impact if recommended shipping conditions and holding times in referenced methods are maintained. Improper packing materials, bulks shipped w/samples, etc. may adversely affect data. Field blanks, field spiked samples or duplicate field samples shipped with samples or included in method validation studies may reflect these contributors. Field variability (FS/DUP) is only considered when lab is responsible for sampling.
lab storage time, conditions and temperature | NA / (LCS, FS) | Usually no impact if recommended storage conditions and holding times in reference methods are maintained. Impact is monitored per sample batch only if LCS samples are prepared on receipt and stored with field samples. Field spiked samples or stability study samples included in some method validation studies may also reflect these contributors. Field variability (FS/DUP) is only considered when lab is responsible for sampling.
contamination in lab storage areas | NA / LCS, FB | Usually no impact if recommended storage conditions and holding times in reference methods are maintained. Improper storage such as sorbent tubes stored with bulk solvent samples or near solvent sources may adversely affect data. Impact is monitored only if LCS is prepared on receipt & stored with field samples. Field blank can be used to assess contamination from collection, transport, and storage.
Laboratory Subsampling |  | NOTE: This primarily applies to bulk/solid samples which requires use of laboratory duplicate sample data to determine impact on uncertainty. It is not applicable for LCS/LCSD media spike duplicates.
sample nonhomogeneity | DUP | Sample composition, etc.
blending techniques | DUP | Stirring, sieving, grinding, etc.
sample size | DUP | Large enough to allow adequate subsampling.
Sample Preparation:
volumetric glassware | LCS, DUP | NA for Class A; applies for graduated tubes or cylinders, etc.
dispensing device | LCS, DUP | Pipettes, and other types of dispensers that are not Class A.
balance | LCS, DUP | Balance error is often insignificant compared to other MU sources.
temperature | LCS, DUP | Hot plate or ashing temperatures.
sample extraction | LCS, DUP | Applies to LCS or DUP if it goes throughthe entire sample preparation process.
extractant background | LCS, DUP, MB | Analyte or interferant is present in acids, solvents, etc.
Lab Environmental Conditions:
temperature variance | LCS, DUP | Room temp during bulk asbestos, gravimetry, etc. processes.
humidity variance | LCS, DUP | Gravimetry involving hydroscopic media, etc.
Analysts:
different analysts | LCS, DUP | Must use inter-analyst instead of intra-analyst repeat data, where applicable.
analyst training level and experience | LCS, DUP | Must use inter-analyst instead of intra-analyst repeat data, where applicable.
data interpretation by analyst | LCS, DUP | Chromatographic peak ID, interference corrections, etc. Must use inter-analyst instead of intra-analyst repeat data, where applicable.
Measuring Instruments:
instrument stability | LCS, DUP | Baseline drift, repeatability of averaged readings, lab environmental stability, etc.
carry over effects | LCS, DUP | Impact of high samples on following sample readings; can be monitored by proper use of CCBs.
day to day calibration differences | LCS | Variation in instrument response and calibration process
interferences | LCS | Due to matrix, inter-element effects, co-eluting GC peaks, etc.
Calibration Standards/Reference Materials:
preparation variances | LCS | Due to analysts, balances, dispensing devices used, etc
calibration stock material uncertainty | CERTIFICATE | Obtain from certificate or estimate, can be ignored if less than 1/3 of the largest contributor.
LCS reference material uncertainty | CERTIFICATE | Only has impact when LCS data are used to correct customer sample results. Obtain from certificate or estimate. Can be ignored if less than 1/3 of the largest contributor. Note that use of an LCS with a large uncertainty can result in over estimation of overall analytical uncertainty.
Test Procedure Variations:
variation within and between reagent lots | LCS | Similar to extractant background effects under Sample Preparation above.
extraction or digestion times, temperatures, and conditions | LCS | May affect complete dissolution of analyte or loss of material in some cases.
sample dependent modifications | LCS | Changes in conditions due to sample size, customer requests, etc.
desorption efficiencies within and between lots for sorbents | LCS | May vary by lot or manufacturer; also applies to diffusion rates for passive monitors.
Data Manipulation:
sampling media/blank correction | LCS, MB | When significant and when data are blank corrected.
instrument blank correction | LCS | When allowed.
accuracy of calculations | LCS | Manual, spreadsheet, LIMS, etc.
area or air volume sampled | NA | Typically provided by the customer. This is not part of analytical uncertainty, but must be considered by labs providing sampling and providing combined sampling and analytical uncertainty.
Where:
DUP = Duplicate, resulting from sub-sampling of a bulk (NOTE: NOT LCS/LCSD duplicate spiked sampling media)
FB = Field Blank
FS = Field Spike - Not typically conducted unless part of sampling method validation. Should be considered only when laboratory is responsible for field sampling.
LCS = Laboratory Control Standard, matrix matched and typically taken through the entire analytical process with each sample batch
MB = Method or Matrix Blank
NA = Not Applicable

## Chromatography Example
Example of Contributors to Measurement Uncertainty Chemical Analyses of Sorbent Tubes using Chromatography See Example Calculations (to the right of the table) |  |  | Examples of Analytical Measurement Uncertainty for Analysis of Organic Solvents on Charcoal Tubes
|  |  | mg Toluene on Charcoal Tube using CS2 Desorption per NIOSH 1501 LCS = 0.0867 mg toluene spike/tube |  |  |  | mg n-Butanol on Charcoal Tube using CS2 Desorption per NIOSH 1401Mod LCS = 0.0810 mg n-butanol spike /tube (includes correction for DE = 0.80)
Contributors to Uncertainty | Representative and Applicable QC Data | Comments to Clarify Contributor Effects
Transportation/Storage/Handling |  |  |  | LCS mg | LCS %Rec |  |  | LCS mg | LCS %Rec
shipping time, container & temperature | NA | Limited impact on most sorbent tubes |  | 0.0876 | 101.0381 |  |  | 0.0862 | 106.4198
lab storage time, conditions & temperature | NA | Usually no impact if recommended storage conditions and hold times are maintained. LCS samples are representative if prepared on receipt & stored with field samples - usually no impact if recommended storage conditions and holding times are maintained. Field spiked samples or stability study samples included in some method validation studies may also reflect these contributors. |  | 0.0824 | 95.0404 |  |  | 0.0799 | 98.642
contamination in lab storage areas | NA | Usually no impact if appropriate storage conditions are maintained. Field blank can be used to assess contamination from collection, transport, and storage |  | 0.0867 | 100 |  |  | 0.0839 | 103.5802
Laboratory Subsampling |  |  |  | 0.0832 | 95.9631 |  |  | 0.0923 | 113.9506
sample nonhomogeneity | NA | Not applicable to sorbent tube analysis |  | 0.0789 | 91.0035 |  |  | 0.0847 | 104.5679
blending techniques | NA | Not applicable to sorbent tube analysis |  | 0.0798 | 92.0415 |  |  | 0.0702 | 86.6667
sample size | NA | Not applicable to sorbent tube analysis |  | 0.085 | 98.0392 |  |  | 0.0632 | 78.0247
Sample Preparation: |  |  |  | 0.0824 | 95.0404 |  |  | 0.0933 | 115.1852
volumetric glassware | LCS | Same type of glassware used for samples and LCS |  | 0.0858 | 98.9619 |  |  | 0.1005 | 124.0741
dispensing device | LCS | Same type of dispensing device |  | 0.085 | 98.0392 |  |  | 0.0987 | 121.8519
balance | NA | Not applicable to sorbent tube analysis |  | 0.0876 | 101.0381 |  |  | 0.0826 | 101.9753
temperature | NA | Not applicable to sorbent tube analysis |  | 0.0832 | 95.9631 |  |  | 0.0913 | 112.716
sample extraction | LCS | Applies to LCS if goes through sample preparation |  | 0.0884 | 101.9608 |  |  | 0.077 | 95.0617
extractant background | LCS, MB | Analyte or interferant in solvents or other prep reagents used, etc. |  | 0.0867 | 100 |  |  | 0.091 | 112.3457
Lab Environmental Conditions: |  |  |  | 0.0902 | 104.0369 |  |  | 0.0832 | 102.716
temperature variance | LCS | LCS results reflect any temperature effects on chromatography instrument |  | 0.085 | 98.0392 |  |  | 0.0984 | 121.4815
humidity variance | NA | Not applicable to sorbent tube analysis |  | 0.0876 | 101.0381 |  |  | 0.0872 | 107.6543
Analysts: |  |  |  | 0.0834 | 96.1938 |  |  | 0.0809 | 99.8765
different analysts | LCS | LCS results reflect variability due to different analysts, as applicable, on different days |  | 0.0819 | 94.4637 |  |  | 0.0926 | 114.321
analyst training level & experience | LCS | LCS results reflect variability due to different analysts, as applicable, on different days |  | 0.0797 | 91.9262 |  |  | 0.0979 | 120.8642
data interpretation by analyst | LCS | LCS results reflect variability due to different analysts, as applicable, on different days
Measuring Instruments: |  |  | 20 point Mean % Rec |  | 97.4913 |  | 20 point Mean % Rec |  | 107.0988
instrument stability | LCS | LCS results reflect instrument variability on different days | 20 point Std Dev |  | 3.602 |  | 20 point Std Dev |  | 11.9695
carry over effects | LCS | LCS results reflect instrument variability on different days |  | RSD | 3.695% |  |  | RSD | 11.18%
day to day calibration differences | LCS | LCS results reflect instrument variability on different days | Reference material used for calibration indicates assay of 99.5%. Expanded uncertainty of reference material estimated at 0.5% (95%) Divide by 2 to yield 0.25 % Rel. SD. Insignificant compared to 3.7% can be eliminated from calculation |  |  |  | Reference materials used for calibration and LCS preparation indicate assay of 99.5%. Expanded uncertainty of reference materials estimated at 0.5% (95%). Divide by 2 to yield 0.25 % Rel. SD. Insignificant compared to 11.2% can be eliminated from calculation
interferences | LCS | LCS results reflect instrument variability on different days
Calibration Standards/Reference Materials:
preparation variances | LCS | Due to analysts, dispensing devices used, etc | Expanded MU @ 95% Conf k=2 |  | 7.389% | % | Expanded MU @ 95% Conf k=2 |  | 22.35% | %
calibration stock material uncertainty | CERTIFICATE | Obtain from certificate or estimate | Bias @ 97.5% Rec = -2.5% |  |  |  | Bias @ 107.1% Rec = 7.1%
LCS reference material uncertainty | CERTIFICATE | Use if customer sample data corrected for desorption efficiency. Obtain from certificate or estimate. | Example of reporting for 0.085 mg toluene: 0.085 mg toluene with an analytical uncertainty of +/- 0.006 mg at the 95% confidence level (k=2) with probable bias of -0.002 mg |  |  |  | Example of reporting for 0.085 mg n=butanol: 0.085 mg n-butanol with an analytical uncertainty of +/- 0.019 mg at the 95% confidence level (k=2) with probable bias of 0.006 mg
Test Procedure Variations
variation within and between reagent lots | LCS | LCS subjected to same treatment as customer samples
extraction or digestion times and temps | LCS | LCS subjected to same treatment as customer samples
sample dependent modifications | LCS | LCS subjected to same treatment as customer samples
desorption efficiencies within and between lots for sorbent tubes | LCS | LCS subjected to same treatment as customer samples | No significant background in method blank. |  |  |  | No significant background in method blank.
Data Manipulation:
sampling media blank correction | LCS | LCS subjected to same treatment as customer samples
instrument blank correction | LCS | LCS subjected to same treatment as customer samples
Accuracy of calculations | LCS | LCS subjected to same treatment as customer samples
LCS = Laboratory Control Standard, typically taken through the entire analytical process with each sample batch
MB = Method or matrix blank
NA = Not Applicable

## Spectrometry Example
Examples of Analytical Measurement Uncertainty for Metals in Air
Milligrams Metal on MCE Air Filter using hot plate digestion and ICP-AES in accordance with OSHA ID-125G
LCS Concentrations = 0.062 mg (for Cr or Pb), 0.125 mg Cd, or 0.0125 mg Be metal spike/air filter--Results in % recovery
LCS, mg Cd/sample | Cadmium | LCS, mg Cr/sample | Chromium | LCS, mg Pb/sample | Lead | LCS, mg Be/sample | Beryllium
0.127 | 101.6 | 0.0583 | 94.0323 | 0.062 | 100 | 0.0129 | 103.2
0.126 | 100.8 | 0.058 | 93.5484 | 0.0598 | 96.4516 | 0.0129 | 103.2
0.13 | 104 | 0.0586 | 94.5161 | 0.0623 | 100.4839 | 0.0126 | 100.8
0.132 | 105.6 | 0.0593 | 95.6452 | 0.0619 | 99.8387 | 0.0117 | 93.6
0.129 | 103.2 | 0.0613 | 98.871 | 0.063 | 101.6129 | 0.0121 | 96.8
0.128 | 102.4 | 0.0608 | 98.0645 | 0.0652 | 105.1613 | 0.012 | 96
0.126 | 100.8 | 0.0605 | 97.5806 | 0.0616 | 99.3548 | 0.0128 | 102.4
0.127 | 101.6 | 0.0613 | 98.871 | 0.0591 | 95.3226 | 0.0122 | 97.6
0.12 | 96 | 0.0567 | 91.4516 | 0.0667 | 107.5806 | 0.013 | 104
0.121 | 96.8 | 0.0559 | 90.1613 | 0.0613 | 98.871 | 0.013 | 104
0.123 | 98.4 | 0.0598 | 96.4516 | 0.0592 | 95.4839 | 0.0129 | 103.2
0.124 | 99.2 | 0.0585 | 94.3548 | 0.0624 | 100.6452 | 0.0124 | 99.2
0.122 | 97.6 | 0.0618 | 99.6774 | 0.0588 | 94.8387 | 0.0131 | 104.8
0.13 | 104 | 0.0607 | 97.9032 | 0.0601 | 96.9355 | 0.0128 | 102.4
0.129 | 103.2 | 0.0627 | 101.129 | 0.062 | 100 | 0.0123 | 98.4
0.129 | 103.2 | 0.062 | 100 | 0.0596 | 96.129 | 0.0129 | 103.2
Mean | 101.15 |  | 96.3911 |  | 99.2944 |  | 100.8
Std. Dev. | 2.8317 |  | 3.1639 |  | 3.5307 |  | 3.4067
RSD, % | 2.7995 |  | 3.2823 |  | 3.5558 |  | 3.3796
Expanded MU, % | 5.5991 |  | 6.5646 |  | 7.1116 |  | 6.7592
Bias, % | 1.15 |  | -3.6089 |  | -0.70565 |  | 0.8
Reference material used for calculation and LCS preparation indicates concentration and expanded uncertainty of 1000 ± 3 μg/mL
at 95% confidence interval. Expanded uncertainty divided by 2 to yield 0.15% RSD.
Insignificant compared to the relative standard deviations calculated above (range 2.8% to 3.6%) and can be eliminated from the calculation.
No significant background in method blank.
Reporting example for 0.125 mg cadmium:
0.125 mg cadmium with an analytical uncertainty of ± 0.007 mg at the 95% confidence level (k=2) with a probable bias of 0.001 mg

## Pb in Paint Example
Example of Contributors to Measurement Uncertainty Chemical Analyses of Lead (Pb) in Paint using ICP-AES See Example Calculations (to the right of the table) |  |  | Examples of Analytical Measurement Uncertainty for Lead in Paint
|  |  | Pb in Paint using hotblock acid digestion and ICP-AES Analysis in accordance with EPA SW846 3050 & 6010 (Mod) LCS Recovery of Paint SRM 2582 at 208.8 +/- 4.9 mg/Kg or SRM 2581 at 4490 +/- 110 mg/Kg Pb | [5 empty] | Sample duplicata data in mg/Kg for Pb in Paint using hotblock acid digestion and ICP-AES in accordance with EPA SW-846 3050 & 6010 (Mod)
Contributors to Uncertainty | Representative and Applicable QC Data | Comments to Clarify Contributor Effects
Transportation/Storage/Handling
shipping time, container & temperature | NA | No impact on bulk paint samples from transportation, storage or normal handling |  | LCS mg/Kg | True value mg/Kg | LCS % Rec |  |  | mg/Kg Sample | mg/Kg Sple Dup | Std Dev (S) | CV | CV2
lab storage time, conditions & temperature | NA |  |  | 195.4 | 208.8 | 93.5824 |  |  | 1604 | 1502 | 72.1249 | 0.046442 | 0.0021569
contamination in lab storage areas | NA |  |  | 189.6 | 208.8 | 90.8046 |  |  | 511 | 602 | 64.3467 | 0.11563 | 0.01337
Laboratory Subsampling |  |  |  | 4157 | 4490 | 92.5835 |  |  | 9470 | 8794 | 478.0042 | 0.052344 | 0.0027399
sample nonhomogeneity | DUP | Sample composition, etc. |  | 4186 | 4490 | 93.2294 |  |  | 161 | 93 | 48.0833 | 0.37861 | 0.14334
blending techniques | DUP | Stirring, sieving, grinding, etc |  | 184 | 208.8 | 88.1226 |  |  | 1683 | 1411 | 192.333 | 0.12433 | 0.015457
sample size | DUP | Large enough to allow adequate subsampling |  | 184 | 208.8 | 88.1226 |  |  | 956 | 830 | 89.0955 | 0.099771 | 0.0099542
Sample Preparation: |  |  |  | 4641 | 4490 | 103.363 |  |  | 23470 | 26570 | 2192.031 | 0.087611 | 0.0076757
volumetric glassware | LCS, DUP | NA for Class A; applies for graduated tubes or cylinders, etc. |  | 4831 | 4490 | 107.5947 |  |  | 625 | 730 | 74.2462 | 0.10959 | 0.01201
dispensing device | LCS, DUP | pipettes, and other types of dispensers not Class A |  | 179 | 208.8 | 85.728 |  |  | 723 | 472 | 177.4838 | 0.29704 | 0.088235
balance | LCS, DUP | balance error is often insignificant compared to other MU sources |  | 177 | 208.8 | 84.7701 |  |  | 23000 | 22000 | 707.1068 | 0.031427 | 0.00098765
temperature | LCS, DUP | Hot plate or ashing temperatures |  | 4539 | 4490 | 101.0913 |  |  | 14190 | 13900 | 205.061 | 0.0146 | 0.00021317
sample extraction | LCS, DUP | Applies to LCS or DUP if goes through sample preparation |  | 4858 | 4490 | 108.196 |  |  | 10350 | 9142 | 854.185 | 0.087645 | 0.0076816
extractant background | LCS, DUP, MB | Analyte or interferant in acids, or other reagents |  | 198 | 208.8 | 94.8276 |  |  | 5702 | 5854 | 107.4802 | 0.018602 | 0.00034602
Lab Environmental Conditions: |  |  |  | 207 | 208.8 | 99.1379 |  |  | 109 | 122 | 9.1924 | 0.079588 | 0.0063342
temperature variance | NA | No impact on bulk paint samples |  | 4458 | 4490 | 99.2873 |  |  | 7079 | 6427 | 461.0336 | 0.068271 | 0.0046609
humidity variance | NA | No impact on bulk paint samples |  | 4514 | 4490 | 100.5345 |  |  | 196 | 186 | 7.0711 | 0.037021 | 0.0013706
Analysts: |  | Analyst contributors affect all aspects of analysis from subsampling through data manipulation |  | 184 | 208.8 | 88.1226 |  |  | 14510 | 14300 | 148.4924 | 0.010308 | 0.00010626
different analysts | LCS, DUP |  |  | 187 | 208.8 | 89.5594 |  |  | 15710 | 17150 | 1018.2338 | 0.061974 | 0.0038408
analyst training level & experience | LCS, DUP |  |  | 4551 | 4490 | 101.3586 |  |  | 1187 | 1192 | 3.5355 | 0.0029723 | 8.8345e-06
data interpretation by analyst | LCS, DUP |  |  | 4580 | 4490 | 102.0045 |  |  | 9265 | 9246 | 13.435 | 0.0014516 | 2.1071e-06
Measuring Instruments: |  |  |  | 179 | 208.8 | 85.728 |  |  | 4240 | 3918 | 227.6884 | 0.05582 | 0.0031158
instrument stability | LCS | Baseline drift, repeatability of averaged readings, etc |  | 186 | 208.8 | 89.0805 |  |  | 1979 | 1574 | 286.3782 | 0.1612 | 0.025987
carry over effects | LCS, DUP | Impact of high samples on following sample readings; can be monitored by proper use of CCBs |  | 4245 | 4490 | 94.5434 |  |  | 2357 | 3068 | 502.7529 | 0.18535 | 0.034353
day to day calibration differences | LCS |  |  | 4302 | 4490 | 95.8129 |  |  | 2254 | 3062 | 571.3423 | 0.21495 | 0.046204
interferences | DUP, MS | Due to matrix, inter-element effects, etc. Cannot be routinely determined for typical industrial hygiene sampling media |  | 188 | 208.8 | 90.0383 |  |  | 53160 | 44300 | 6264.9661 | 0.12856 | 0.016529
Calibration Standards/Reference Materials: |  |  |  | 202 | 208.8 | 96.7433 |  |  | 24810 | 23000 | 1279.8633 | 0.05354 | 0.0028665
preparation variances | LCS, DUP | Due to analysts, balances, dispensing devices used, etc |  | 4217 | 4490 | 93.9198 |  |  | 22860 | 22930 | 49.4975 | 0.0021619 | 4.674e-06
calibration stock material uncertainty | CERTIFICATE | Obtain from certificate or estimate |  | 4300 | 4490 | 95.7684 |  |  | 1133 | 1125 | 5.6569 | 0.0050105 | 2.5105e-05
LCS reference material uncertainty | NA | Sample results not corrected for LCS recovery |  | 203 | 208.8 | 97.2222 |  |  | 151 | 96 | 38.8909 | 0.31491 | 0.099166
Test Procedure Variations |  |  |  | 190 | 208.8 | 90.9962 |  |  | 7774 | 4366 | 2409.8199 | 0.397 | 0.15761
variation within and between reagent lots | LCS | Similar to extractant background effects under Sample Preparation above |  |  | 30 point Mean % Rec. | 94.7291 | [5 empty] | ∑ CV2 | 0.70636
extraction or digestion times and temps | LCS | May affect complete dissolution of analyte or loss of material in some cases | 30 point Std Dev |  |  | 6.3143 |  |  |  | CV pooled = √ (∑CV2/30) = |  |  | 0.15344 | 15.3% | RSD
sample dependent modifications | LCS | Changes in conditions due to sample size, customer requests, etc |  |  | RSD | 6.666%
desorption efficiencies within and between lots for sorbent tubes | NA |  |  |  |  | Reference material used for calibration indicates concentration and expanded uncertainty of 1001 +/- 5 ug/mL at 95% confidence level. Expanded uncertainty divided by 2 to yield 0.25 % Rel. SD. Insignificant compared to 15.2% and can be eliminated from calculation. No significant background in method blank.
Data Manipulation:
sampling media blank correction | NA | No sampling media with bulk samples
instrument blank correction | LCS | when allowed
Accuracy of calculations | LCS | Manual, spreadsheet, LIMS, etc |  |  |  |  | Combined Rel. Std Dev (SDc) = √ [ SD12 + SD22]
SDc = √[(6.7)2 + (15.3)2] = 16.7%
[7 empty] | Expanded MU @ 95% Conf (k=2) = |  |  |  | 33.4%
Bias @ 94.7% Rec of LCS = -5.3%
DUP = Duplicate, resulting from sub-sampling of a bulk (NOTE: NOT LCS/LCSD duplicate spiked sampling media)
FB = Field Blank |  |  |  |  | Example analytical uncertainty for 4400 mg/Kg Pb in paint sample: Expanded analytical uncertainty of 4400 mg/Kg Pb in paint = 4400 X 0.334 = 1470 mg/kg
FS = Field Spike
LCS = Laboratory Control Standard, matrix matched and typically taken through the entire analytical process, with each sample batch
MB = Method or matrix blank |  |  |  |  | Bias = 4400 mg/kg X -0.053 = 233.2 mg/kg
NA = Not Applicable
Example of reporting for 4400 mg/Kg Pb in paint: 4400 mg/Kg Pb in paint with an analytical uncertainty of +/- 1500 mg/Kg at the 95% confidence level (k=2) and a probable bias of -230 mg/Kg